# DropBot self test (2017-09-12T19:19:36)

# Control board (UUID: 012345678-abcd-ef01-234-56789abcdef0)

## Properties

- display_name: DropBot

- package_name: dropbot

- software_version: 1.30

- url: http://gitlab.com/sci-bots/dropbot.py.git

- base_node_software_version: 0.29

- manufacturer: Sci-Bots Inc.

## Config

- max_voltage: 150.0

- R7: 10984.2734375

- min_frequency: 100.0

- switching_board_i2c_address: 32

- max_frequency: 10000.0

- pot_max: 49657.4570312

- i2c_address: 10

- id: ``

soft i2c scan: [44]
number of channels: 120

# I2C scan:

- 32: HV Switching Board v3.1, firmware: v0.5, uuid: 012345678-abcd-ef01-234-56789abcdef1

- 33: HV Switching Board v3.1, firmware: v0.5, uuid: 012345678-abcd-ef01-234-56789abcdef2

- 34: HV Switching Board v3.1, firmware: v0.5, uuid: 012345678-abcd-ef01-234-56789abcdef3

- 80: 120-channel-pogo-pin-board v3.3, uuid: 012345678-abcd-ef01-234-56789abcdef4

# Test voltage results:

- Output voltages:

- 0       1       2        3        4
target    46.2 V  67.0 V  87.7 V  108.5 V  129.2 V
measured  46.2 V  66.7 V  87.4 V  109.5 V  131.3 V

- Root-mean-squared (RMS) error: 0.9%

# Test shorts results:

- No shorts.

# Test on-board feedback calibration results:

- Measured capacitance:

- 0        1         2         3
nominal    0.0 F  10.0 pF  100.0 pF  470.0 pF
measured  2.0 pF  10.4 pF  111.2 pF  494.7 pF

- On-board feedback calibration capacitors

# Test channels results:

All 120 channels passed.

Channel capacitance summary